# Dynamic ISF — can sensitivity be derived from data? (feasibility findings)

2026-06-07 · Tim Street / Claude · companion to the methodology and v1/v2 analysis

## Question

Can a user’s insulin sensitivity (ISF) be derived from device-observable data well enough to remove the need to enter it — and, if so, how? This summarises a six-phase feasibility study on 138 open-source AID users (Trio, AAPS, OpenAPS), reusing the empirical-ISF ΔIOB estimator and the √TDD law established in the equation-fit analysis.

## The two separable parts of ISF

ISF has a shape, meaning how it scales with TDD and glucose, and a per-user level. The two behave very differently.

The shape is derivable and settled. Sensitivity scales with total daily dose as roughly 1/√TDD, and that holds across users, with the glucose curve a separate dimension on top. The shape needs no input from anyone.

The level is the hard part, and the subject of this study. A single global level leaves the typical person about 57% off, so the question is whether the per-user level can be recovered from data rather than entered.

## The feasibility arc (Phases 1–6)

| phase | question | finding |
| --- | --- | --- |
| 1 | Is a short-window empirical ISF dosing-grade? | No. ±10% CI only at 60–90 d; a 14-day estimate swings ±34% block-to-block, memoryless (lag-1 ACF ≈ 0). |
| 2 | Is the variation trackable structure? | Mostly not. Circadian ~2–4%, weekly ~1%, slow ~2%; ~90% residual, near-white, partly measurement noise. |
| 3 | Can a stable baseline be recovered by shrinkage? | Yes, in self-consistency terms: cold-start ±42%; shrinking own-data toward the √TDD prior → ±22% (14 d) → ±13% (90 d). |
| 5 | Does the derived (sensitivity-anchored) level beat tuned ISF on real outcomes? | No — and it’s unsafe. Users dosing far “weaker” than their measured sensitivity have more hypoglycaemia, not less (logR-vs-TBR +0.38). The empirical ISF is hypo-biased. |
| 6 | Does overnight data de-bias the sensitivity estimate? | No. Overnight is noisier (±50% test-retest) and carries the same bias (+0.39). The bias is intrinsic to observational sensitivity, not a daytime artefact. |
| 7 | Can outcome-based tuning give a safe per-user level? | Yes, in principle. Working-anchored cold start (355/√TDD) lands ±28% of the users’ tuned ISF with no entry; a bounded outcome nudge removes the residual hypo-signature (logR-vs-TBR +0.21 → −0.04) where the sensitivity route made it worse (+0.38). |

## What this settles

- The √TDD shape is derivable and needs no entry. Anchored to working (tuned) profiles the constant is K ≈ 355 (ISF ≈ 355/√TDD); anchored to measured sensitivity it is ≈ 145. Since Phases 5 and 6 show the measured anchor is unsafe, the working-anchored 355 is the right cold start.

- A zero-entry cold start is feasible and reasonably good. From observable TDD alone, ISF ≈ 355/√TDD places a brand-new user within ±28% of the value they would have tuned to (Phase 7). Against the measured empirical baseline the figure is ±42%, but that baseline is the hypo-biased one, so the working-ISF comparison is the number that matters for deployment.

- The per-user level cannot be safely derived from observed sensitivity, in any window. The ΔIOB sensitivity estimate is entangled with the user’s existing insulin excess or deficit: it reads as “very sensitive” (low ISF) precisely for people who run low, so dosing to it would give the most insulin to the most hypo-prone. Restricting to clean overnight data does not fix this. It is intrinsic to measuring sensitivity from observational glucose against insulin.

- Beyond a stable baseline there is little left to track for correction ISF (Phase 2). The right architecture is therefore not an elaborate adaptive ISF but a stable baseline plus the √TDD demand-track, with autosens covering the bounded sub-daily and safety role it already serves.

## The constructive conclusion

A data-derived, zero-entry ISF is feasible only in this form:

√TDD shape (universal) + working-anchored cold start (≈ 355/√TDD, ±42%, no entry) + bounded outcome-based self-tuning of the level — nudging ISF weaker when unexplained lows appear and stronger when sustained unexplained highs appear, damped and clamped.

It is not feasible as “measure the user’s sensitivity and dose to it” — that route is hypo-biased regardless of how clean the window is. The level must be tuned from outcomes (the autotune/autosens philosophy), not from a sensitivity regression. That is the autotune design, now empirically explained: it is built on the human-tuned anchor and outcome-attribution precisely because the measurement route fails.

## Caveats

Decision-level and observational (no closed-loop simulation); single cohort, mostly users who never ran dynamic ISF; the empirical estimator carries carb/endogenous-glucose confounds; “level” errors in Phase 3 are self-consistency, not outcome-validated correctness. A closed-loop trial is required before any of this drives dosing. Nothing here is medical advice.

## Reproducibility

inv008/phase{1,2,3,5,6}_*.py and results/phase*_*.{json,md} in github.com/tim2000s/dynamic-isf-calculations. Inputs device-observable; the user’s profile ISF is used as a benchmark only.

## Addendum, 27 August 2026: replication on independent data

Phases 5 and 6 above concluded that observational sensitivity is biased low, that dosing to it is unsafe, and that the bias is intrinsic rather than a daytime artefact. Those conclusions rested on 138 users of open-source systems with insulin on board taken from the loop’s own record.

They have since been tested on 1,679 people across six Jaeb trial cohorts, ages 2 to 82, with insulin on board reconstructed from delivery records rather than reported, and using four different constructions of the estimator. The ratio of measured to entered sensitivity ranges from 0.30 to 0.92 with a median of 0.58, and in fifteen estimates it does not once reach 1.0. The K = 145 against 355 reported here corresponds to a ratio of 0.41 and sits inside that range.

So the finding is not a property of this cohort, of open-source systems, or of how insulin on board was obtained. See Dynamic-ISF-7.